## System Analysis
MIT-E VENT. TopDrive System Design spreadsheet
Spreadsheet to deterministically design ambu bag squeezer using pure rotary linkage with sontinuous rotary motion input. COVID sadness can bring world together to work as a team, and then maybe next step fix climate change, hence one over sadness equals happiness 1/:-( = :-). see www.ooseh.com
created by Prof. Alex Slocum March 21, 2020, see "Revisions" tab
Inputs in BLACK, Outputs in RED
Respiration |  |  |  | Comments
Respiration rate
Peak for shorter periods | resprate | 35 | breaths per minute
Average for mechanism life calculations | resprateavg | 30 | breaths per minute
arc of travel of finger to compress bag | arcfing | 25 | degrees
seconds per breath | spb | 1.7143
Inspiratory/expiratory ratio | ier | 1:4
Inspiration ratio | insp | 1 | portion of cycle
Expiration ratio | expr | 4 | portion of cycle
Total parts | tpb | 5 | total portions
Time to inhale | tinsp | 0.34286 | seconds
Time to exhale | texh | 1.3714 | seconds
Portion of inhalation to ramp | pcr | 0.1 | percentage
time to ramp to pressure | trtp | 0.034286 | seconds
effective average rpm of finger | rpmfing | 12.1528 | rpm
Drive transmission ratio (enter gear sizes below) | TRpf | 3
required instantaneous rpm of gearmotor shaft | rpmgmr | 36.4583 | rpm | : =trpf*Rpmfing
Forces on "Hands" from airway pressure in bag
peak airway pressure | pap | 50 | mmHg
|  | 65 | cmH2O
factor higher to force from bag through circuit | papeta | 0.9
bag pressure | pbagmm | 55.5556 | mmHg
|  | 0.073099 | atm
|  | 1.0746 | psi
|  | 0.0074118 | N/mm^2
| cmH2O | 72.2222 | cmH2O
Bag diameter | Dbag | 125 | mm
| Lbag | 125 | mm
| Areabag | 15625 | mm^2
force on bag | Fbag | 115.8094 | N
|  | 26.0246 | lbs | force on EACH finger
Estimated Motor Size Required for the System
Life cycle
Breaths per day | bpdavg | 43200
Days of life of machine before rebuild | dolm | 180
Total machine cycles required | tmcd | 7776000
Motor Power estimate to start design
Radial velocity of bag sides during inhalation | Vrbagsides | 182.2917 | mm/sec | : =Dbag/2/tinsp
Power that must be applied to compress bag during inhalation portion of breathing cycle | Pbaginhale | 42.2222 | Watts | : =2*Fbag*Vrbagsides/1000. NOTE the factor of 2 is because one motor drives both fingers
mechanical system effieicny (gearmotor and connection to bag mechanism) | etamech | 0.8
Oversize motor for 100% duty cycle | ovszf | 1
Estimated total peak motor mechanical power to select candidate motors | Pmotorest | 52.7777 | Watts
Geared Hands Design: MIT-E Vent | [5 empty] | 2
Orientation of the pinion-to-drive gear contact force mid-bag compression: for a "pure" top drive Theta = 0. For a "pure" bottom drive drive Theta = 180. For a "pure" side drive Theta = 90
Theta (angle at which the pinion force engages the finger drive gear) | Thetapingear | 180
distance finger center to finger contact point | R_1 | 139.7 | mm |  |  | #REF!
|  | 5.5 | inch
distance finger center to center plane | R_2 | 38.1 | mm
|  | 1.5 | inch
diameter of driver gear | Dgear | 76.2 | mm
Radius of driver gear | Rgear | 38.1 | mm
|  | 1.5 | inch
diameter of finger support shaft | Dbear | 9.525 | mm
|  | 0.375 | inch
coefficient of friction of finger support bearing | mu | 0.1
Force on a each "hand" as bag is compressed | Ffinger | 115.8094 | N
Force on gear teeth between fingers to share torque from gearmotor | Fgtg | 424.6345 | N
Required motor pinion tooth force on driven gear to move both "Hands" together to compress bag | Fgear | 849.269 | N | The pinion must be able to drive BOTH fingers, hence the factor of "2"
|  | 190.847 | lbf
Torque generated by driven gear which is then shared by both Hands | Tgear | 32357.1486 | N-mm
|  | 286.2704 | lb-in
Hand shaft support Bearing forces
Y direction | Fbx | -965.0784
X direction | Fby | 424.6345
Force on finger support shaft (and bearings) | Fbear | 1054.3675 | N
radial load per bearing | Fperbear | 527.1837 | N
|  | 118.4683 | lb
Torque lost to shaft bearing friction | Tlsbf | 502.1425 | N-mm
% torque lost to shaft bearing friction | ptlsf | 1.552% |  |  | 0
Gearmotor
Pinion diameter (NOTE cannot go smaller than 1" for 12 pitch gear) | Dpinion | 25.4 | mm
|  | 1 | inch
Transmission ratio motor pinion-to-finger | TRpf | 3
gearbox output torque needed | Tgm | 10785.7162 | N-mm
|  | 95.4235 | lb-in
gearbox output speed needed | gosn | 36.4583 | rpm
NOTE if during the inhale time tinhale (e.g. 0.4 seconds total) if assume x% for ramp up and x% for ramp down , then over total travel of Y degrees, the motor speed required (upper limit) is Y/(tinhale*(1-2x))
Percent of inhale time for ramp up and ramp down | xpramp | 10%
Maximum desired gearbox speed needed to enable ramp up | gosnact | 45.5729 | rpm
Motor Selection and operating parameters |  |  |  |  | Motor Data from "Motor Options" workbook
N motor size (Nema 23, 34 or Gearmotor 1, 2,3. see workbook "Motor Options") | Nmsize | 1 |  |  | Motor number | 23 | 34 | 1 | 2 | 3
Stall torque | Tstallsm | 22.5156 | N-mm |  | Stall torque (N-m) | 1.8014 | 4.2033 | 22.5156 | 15.6 | 10
| Tstallsmm | 22515.621 | N-mm
Free speed | wfree34 | 75 | rpm |  | no load speed (rpm) | 900 | 720 | 75 | 59 | 59
max radial force 20mm from face | Fradmaxm | 89 | N |  | allowable side load (N) | 75 | 220 | 89 | 89 | 100
required speed for respiration cycle | rsarc | 36.4583 | rpm
torque available at required speed | NemaTas | 11570.5275 | N-mm
torque required for desired performance | tgm | 10785.7162 | N-mm
|  | 95.4235 | lb-in
Number of motors required to provide required torque | nmr | 1
Pinion tooth force at pitch line at maximum motor torque at required speed | PFPL | 911.0652 | N
|  | 204.7337 | lbf
Pinion tooth force/motor shaft allowable radial load |  | 10.2367
Is radial force on motor shaft too high? |  | force too high?
cm H2O this torque will provide | cmH2Omax | 150.7668 | cm H2O
|  | 115.9745 | mmHg
Pinion tooth force at pitch line at motor torque at required speed for minimum cm H2O | PFPLr | 849.269 | N | : =tgm/(Dpinion/2)/NMr NOTE that the radial force depends on how many motors being used
|  | 190.847 | lbf
Is radial force on motor shaft too high? |  | force too high? |  | at 35 mm HG load is
cm H2O this torque will provide | cmH2O | 72.2222 | cm H2O
|  | 55.5556 | mmHg

## Motor Options
Enter numbers in BLACK, equivelants are calculated in RED, and then these values are transferred to the "System Analysis" spreadsheet
Motor number referenced on "System Analysis" sheet
1 | AndyMark PG71 gearmotor. https://www.andymark.com/products/hex-pg-series-gearboxes-options
| Stall torque | Tstall | 16.6 | lb-ft | 22.5156 | N-m
| Free speed | wfree | 75 | rpm | 1.309 | rad/sec
| estimated allowable side load on shaft | ? No spec found, so estimate | 20 | lb | 89 | N
| est peak power (1/2 stall torque X 1/2 free speed) |  |  |  | 7.3682 | W
2 | "Windsheild wiper motor" e.g., http://schmittongaromarine.com/site/1922schm/Technical_Specifications_for_Heavy_Duty_Wiper_Motors_309,319,_and_329_series.pdf. https://www.oslv.com/Products/Gearmotor-for-windshield-wipers-MRT62-43/98-25-1-en.html
| Stall torque | Tstall | 11.5013 | lb-ft | 15.6 | N-m
| Free speed | wfree | 59 | rpm | 1.0297 | rad/sec
| estimated allowable side load on shaft | ? No spec found, so estimate | 20 | lb | 89 | N
| est peak power (1/2 stall torque X 1/2 free speed) |  |  |  | 4.016 | W
3 | Enter motor info
| Stall torque | Tstall | 7.3727 | lb-ft | 10 | N-m
| Free speed | wfree | 59 | rpm | 1.0297 | rad/sec
| estimated allowable side load on shaft | ? No spec found, so estimate | 22.4719 | lb | 100 | N
| est peak power (1/2 stall torque X 1/2 free speed) |  |  |  | 2.5744 | W
23 | Nema 23 Stepper motor
| Stall torque | Tstall | 255 | oz-in | 1.8014 | N-mm
| Free speed | wfree | 900 | rpm | 15.708 | rad/sec
| estimated allowable side load on shaft |  | 16.8539 | lb | 75 | N
| est peak power (1/2 stall torque X 1/2 free speed) |  |  |  | 7.0741 | W
34 | Nema 34 Stepper motor
| Stall torque | Tstall | 595 | oz-in | 4.2033 | N-m
| Free speed | wfree | 720 | rpm | 12.5664 | rad/sec
| estimated allowable side load on shaft |  | 49.4382 | lb | 220 | N
| est peak power (1/2 stall torque X 1/2 free speed) |  |  |  | 13.2051 | W
https://szbobet.en.alibaba.com/product/60516551996-803622807/BM_Brand_High_precision_4_Axis_8_5Nm_1204oz_in_Nema_34_Stepper_Motor_with_braker_Driver_Kit_For_CNC_Router.html
https://hardware-cnc.nl/en/shop/stappenmotoren/stepper-motor-34/stepper-motor-nema-34-4,5nm-detail
NEMA 23
https://cnc1.lv/en/stepper-motors/48-nema-23-stepper-motor-23hs9430-with-280-nm.html

## Gears
Spur_Gears.xls
Spreadsheet to estimate gear tooth strength
Production gears must be designed using the Lewis Form Factor or FEA
By Alex Slocum 1/18/01, last modified 9/28/2007 by Alex Slocum
Enters numbers in BOLD, Results in RED
Inputs
| Motor Pinion | finger to finger
Torque, T (in-lb) | 95.4235 | 143.1352
N-mm | 10785.7162 | 16178.5743
Pressure angle, f (deg, rad) | 14.5 | 0.25307
Pitch, P | 12 | 12
Number of teeth on pinion, Np | 12 | 30
Number of teeth on gear, Ng | 96 | 30
Center distance tolerance, Ctol (inches) | 0.005 | 0.005
Face width (of drive pinion), w (inches) | 0.75 | 3
Pinion material | Steel | Acetal (Delrin)
Pinion yeild stress, sigp (psi) | 40000 | 11500
Gear Material | Acetal (Delrin)
Gear yield stress, sigg (psi) | 11500 | 11500 | Delrin (Acetal). http://www.sdplastics.com/delrin/delrin[1].pdf
Stress concentration factor at tooth root, scf | 1.5 | 1.5
Outputs
Gear ratio, mg | 8 | 1 | : =Ng/Np
Pinion pitch diameter, Dp (inches) | 1 | 2.5 | : =Np/P
Gear pich diameter, Dg (inches) | 8 | 2.5 | : =Ng/P
Center distance, C (inches) | 4.505 | 2.505 | : =0.5*(Dp+Dg)+Ctol
Tooth thickness, tt (inches) | 0.1309 | 0.1309 | : =0.5*PI()/P
Addendum, a (inches) | 0.083333 | 0.083333 | : =1/P
Dedendum, b (inches) | 0.102 | 0.102 | : =1.2/P+0.002
Clearance, cl (inches) | 0.018667 | 0.018667 | :=0.2/P+0.002
Pinion tooth force (normal to radius), Fp (lbf) | 190.847 | 114.5082 | :=T/(Dp/2) |  | 509.5614
Face width assumed to be the molded thickness, force/finger (lbf) | 190.847 | 28.627
Gear tooth separation force (parallel to radius) Fsp (lbf) | 49.3564 | 29.6138 | : =Fp*TAN(f)
Total radial load on motor shaft (lbf) | 197.1259 | 118.2755 | :=SQRT(Fsp^2+Fp^2)
Tooth section parameters | [5 empty] | 0.625 | 0.62432
Chordal area, Ac (inches^2) | 0.098175 | 0.3927 | : =w*tt
First moment of area, Q (inches^3) | 0.0021418 | 0.0085674 | :=w*tt^2/6
2nd moment of area ("moment of inertia"), I (inches^4) | 0.00014018 | 0.00056073 | :=w*tt^3/12
Distance Nuetral axis to outer fiber, cc (inches) | 0.06545 | 0.06545 | :=tt/2
| Motor side gear tooth stress |  | Finger-to-Finger gear stress
Pinion tooth stresses | Stress | Safety Factor | Stress | Safety Factor
Shear stress of the tooth (F/A) (psi) | 1943.9512 | 3.4155 | 291.5927 | 3.4155 | :=Fp/Ac
Bending shear stress (FQ/wI) (psi) | 3887.9024 | 1.7077 | 583.1854 | 1.7077 | :=Fp*Q/(w*I)
Bending stress (F*b*c/I) (psi) | 13632.9365 | 0.84355 | 2044.9405 | 0.48702 | :=Fp*b*cc/I*scf
Von Mises stress (psi) | 14042.5709 | 0.81894 | 2106.3856 | 0.47281 | :=SQRT(3*B39^2+B41^2)
96.8937

## Hex drive
|  | Metal torque arm | plastic fingers
hex shaft diameter | hsd | 0.59055 | 0.59055 | in
|  | 15 | 15 | mm
Width of contact face | wcf | 0.75 | 3 | in
|  | 19.05 | 76.2 | mm
torque to be transmitted | ttbt | 286.2704 | 286.2704 | lb-in
constant from the torque integral | Chexpres | 0.063968
Number of hex faces assumed to transmit the torque | Nhftt | 3
contact pressure on hex faces | hcpf | 22812.6144 | 5703.1536 | psi
|  | 157.3503 | 39.3376 | MPa
NOTE 1/2" hex yields plastic contact pressure of 8,649 psi which is ok for Nylon 6-6 30% glass filled but tie in with steel driver for security
SKF: https://www.skf.com/binary/tcm:12-69175/457649.pdf
Timken: https://www.timken.com/wp-content/uploads/2018/04/Timken-Ball-Bearing-Catalog_10734.pdf

## Bearings
Radial force per bearing | Frbearing | 527.1837 | N
width of bearing | wbearing | 9.525 | mm
diameter bearings | dbearing | 9.525 | mm
contact stress on bearings |  | 5.8107 | N/mm^2
|  | 842.4409 | psi

## lung compliance notes
Lung Compliance Update; Alexander Slocum, Jr, MD PhD; on 3/22/2020; sent to Albert Kwon and Christoph Nabzdyk 3/22/2020
Clinical scenario: 300-400 lb patient who has a lung compliance of 3? Also patient is bucking, having difficulty breathing.
Normal lung compliance, lower bound | LCNL | 50 | cmH2O
Normal lung compliance, upper bound | LCNU | 100 | cmH2O
Change from lower bound
Bad compliance, ARDS | LCA | 3 | cmH2O | 16.6667
Resp rate lower bound |  | 8 | breaths
Resp rate upper bound |  | 40 | breaths
PEEP | PEEP | 15 | cmH2O | This exists everywhere in the breathing curcuit AFTER the valve; so inspiratory pressure of 30 is between bag & valve
Inspiratory pressure | Pins | 30 | cmH2O
Driving pressure (at machine) | Pmach | 15 | cmH2O
Driving pressure (at mouth) | Pmouth | 15 | cmH2O
Anticipated required airway pressure (need torque to generate this much press) | Preq | 30 | max

## refs
https://en.wikipedia.org/wiki/Mechanical_ventilation
https://www.sciencedirect.com/topics/medicine-and-dentistry/mean-airway-pressure
https://www.sciencedirect.com/topics/medicine-and-dentistry/airway-pressure

## Revisions
Version | date | who | edit
1 | 2020-03-24 | AHS Sr | added NEMA 34 drve parameters and cm H2O equivelant
2 | 2020-03-24 | AHS Sr | added motor rpm for breathe cycle
3 | 2020-03-24 | AHS Sr | added select motors and do motor calcs, and added details on gear teeth for both pinion and finger to finger
4 | 2020-03-24 | AHS Sr | add in Albert Kwin and Alex Jr breathing edits
5 | 2020-03-25 | AHS Sr | added radial load additional design info
6 | 2020-03-26 | AHS Sr | added motor references
7 | 2020-03-27 | AHS Sr | generalizing for any design
8 | 2020-03-28 | AHS Sr | generalizing for any design and adding fixturing and hex drive calcs
8 | 2020-03-29 | AHS Sr | generalizing for any design and adding fixturing and hex drive calcs

## Bag fixture by V
Bag1eD | 50 | mm
Bag2eD | 40 | mm
Veeangle | 45 | degrees
deltaCenter | 3.5355 | mm | distance center heigh shifts when supporting Bags of different end diameters in a Vee